## Sheet1
Code | Message | Description | Macro Area | Sub Area | Type | Category | Suggestion | Weight Index
0 | Analyzer / Control Module loss of communication. | A non-recoverable error occurred in the CM/AM Communications Manager (e.g., loss of connection). | SW | SW | Error | Questionable | HW/SW Error | 0.25
5 | Master controller failure. | Failure of LRC or CRC check. | SW | SW | Error | Questionable | HW/SW Error | 1
7 | <Arm> Arm controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | Error | Rare | N/A | 0.5
8 | Sample Arm controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | N/A | Rare | N/A | 0.5
9 | Intermediate Reagent Arm controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | N/A | Rare | N/A | 0.5
10 | Start Reagent Arm controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | N/A | Rare | N/A | 0.5
11 | Cuvettes controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | Error | Rare | N/A | 0.5
12 | Racks controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | Error | Rare | N/A | 0.5
13 | Optical Reading Units controller initialization failure. | One of the AM processors failed to initialize, due to an Invalid Controller ID. | SW | SW | Error | Rare | N/A | 0.5
14 | <Arm> Arm Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | Error | Rare | N/A | 0.5
15 | Master Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | Error | Rare | N/A | 0.5
16 | Cuvettes Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | Error | Rare | N/A | 0.5
17 | Racks Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | Error | Rare | N/A | 0.5
18 | Optical Reading Units Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | Error | Rare | N/A | 0.5
19 | Sample Arm Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | N/A | Rare | N/A | 0.5
20 | Analyzer internal communication error. | A non-recoverable error occurred in the Controller Communications Manager or the low level CAN bus driver (master or controller side. | SW | SW | Error | Questionable | HW/SW Error | 1
21 | Intermediate Reagent Arm Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | N/A | Rare | N/A | 0.5
22 | Start Reagent Arm Controller software error. | Invalid parameter provided in a function call, message queue or inter-task communication mechanism. | SW | SW | N/A | Rare | N/A | 0.5
25 | Master Controller software error. | Memory allocation failed or stack usage exceeds limit. | SW | SW | Error | Rare | N/A | 0.5
30 | Master Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
31 | Cuvettes Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
32 | Racks Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
33 | Optical Reading Units Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
34 | Rack-Cuvette Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
35 | Master Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Questionable | HW/SW Error | 1
36 | Cuvettes Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Rare | N/A | 0.5
37 | Racks Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Rare | N/A | 0.5
38 | Optical Reading Units Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Rare | N/A | 0.5
39 | Rack-Cuvette Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Rare | N/A | 0.5
40 | Master Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | Error | Rare | N/A | 0.5
41 | Cuvettes Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | Error | Rare | N/A | 0.5
42 | Racks Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | Error | Rare | N/A | 0.5
43 | Optical Reading Units Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | Error | Rare | N/A | 0.5
44 | Rack-Cuvette Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | Error | Rare | N/A | 0.5
45 | Master Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
46 | Cuvettes Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
47 | Racks Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
48 | Optical Reading Units Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
49 | Rack-Cuvette Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
50 | Master Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
51 | Cuvettes Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
52 | Racks Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
53 | Optical Reading Units Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
54 | Rack-Cuvette Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
56 | Interrupted communication between the Analyzer and the Control Module. | The CM has lost the connection with the AM module. | SW | SW | Error | Questionable | HW/SW Error | 0.1
57 | Optical Reading Units Controller software error. | One of the wavelengths not used in ORU acquisitions. | SW | SW | Error | Rare | N/A | 0.5
58 | Downloaded controller Software version (<Downloaded File Name>) is different than Software version actually used by AM (<Used File Name>). | Controller Upgrade has failed since downloaded SW version does not match SW version actually used. | SW | SW | Error | Rare | N/A | 0.5
59 | Interface Module upgrade failed. | IM upgrade failed. | SW | SW | Error | Rare | N/A | 0.5
60 | Incompatible IM software version | Incompatible IM software version. | SW | SW | Error | Rare | N/A | 0.5
61 | Sample Arm Controller software error. | Error detected when calling kernel facilities. | SW | SW | Error | Rare | N/A | 0.5
62 | Intermediate Reagent Arm Controller software error. | Error detected when calling kernel facilities. | SW | SW | N/A | Rare | N/A | 0.5
63 | Start Reagent Arm Controller software error. | Error detected when calling kernel facilities. | SW | SW | N/A | Rare | N/A | 0.5
65 | <Arm> Arm Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | Error | Rare | N/A | 0.5
66 | Sample Arm Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | N/A | Rare | N/A | 0.5
67 | Intermediate Reagent Arm Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | N/A | Rare | N/A | 0.5
68 | Start Reagent Arm Controller software error. | Software logic error such as an invalid path in a switch statement or if.else branch. | SW | SW | N/A | Rare | N/A | 0.5
70 | <Arm> Arm Controller software error. | Invalid state for an event, function call, message, etc. | SW | SW | Error | Rare | N/A | 0.5
71 | Sample Arm Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | N/A | Rare | N/A | 0.5
72 | Intermediate Reagent Arm Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | N/A | Rare | N/A | 0.5
73 | Start Reagent Arm Controller software error. | Invalid state for an event, function call, message, etc | SW | SW | N/A | Rare | N/A | 0.5
75 | <Arm> Arm Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | Error | Rare | N/A | 0.5
76 | Sample Arm Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | N/A | Rare | N/A | 0.5
77 | Intermediate Reagent Arm Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | N/A | Rare | N/A | 0.5
78 | Start Reagent Arm Controller software error. | It was detected that there is a potential for a stack problem. | SW | SW | N/A | Rare | N/A | 0.5
80 | <Arm> Arm Controller software error. | The command received is not supported. | SW | SW | Error | Rare | N/A | 0.5
81 | Sample Arm Controller software error. | The command received is not supported. | SW | SW | N/A | Rare | N/A | 0.5
82 | Intermediate Reagent Arm Controller software error. | The command received is not supported. | SW | SW | N/A | Rare | N/A | 0.5
83 | Start Reagent Arm Controller software error. | The command received is not supported. | SW | SW | N/A | Rare | N/A | 0.5
84 | Incompatible AM Master software version. Detected interface version <Interface Version> and expected <Interface Version>. | Incompatible interface version between AM/CM. | SW | SW | Error | Rare | N/A | 0.5
85 | Incompatible <Controller Name> controller version. | Incompatible interface version between a controller and the AM Master software. | SW | SW | Error | Rare | N/A | 0.5
86 | Material <Material Name> not dispensed. | SW detected a material that was not dispensed. | SW | SW | Error | Rare | N/A | 0.5
1001 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1002 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1003 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1004 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1005 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1006 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1007 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1008 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1009 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1010 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1011 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1012 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1013 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1014 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | N/A | Rare | N/A | 0.5
1015 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | N/A | Rare | N/A | 0.5
1016 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | N/A | Rare | N/A | 0.5
1017 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1018 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | N/A | Rare | N/A | 0.5
1019 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | N/A | Rare | N/A | 0.5
1020 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1021 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1022 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1023 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1024 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1025 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1026 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1027 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1028 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Rare | N/A | 0.5
1029 | <Arm> Arm failure. | The CAVRO arm is inoperable | ARMs | Cavro Arms | Error | Questionable | Service Call | 1
1031 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1032 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1033 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1034 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1035 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Questionable | Service Call | 1
1036 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1037 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1038 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Rare | N/A | 0.5
1039 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Questionable | Service Call | 1
1040 | <Syringe> Syringe failure. | The pump used to aspirate/dispense fluids is inoperable. | ARMs | Syringes | Error | Questionable | Service Call | 1
1041 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Warning | Questionable | ORU Calibration | 1
1042 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Error | Rare | N/A | 0.5
1043 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Error | Rare | N/A | 0.5
1044 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Error | Questionable | ORU Calibration | 1
1045 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Error | Questionable | ORU Calibration | 1
1046 | The Optical Reading Unit <ORU ID> has been disabled. | The ORU Head has been disabled. | ORU | Data Check | Error | Rare | N/A | 0.5
1047 | <Syringe> Syringe failure. | Plunger Overload. This error occurs when movement of the syringe plunger is blocked by excessive pressure. | ARMs | Syringes | Error | Rare | N/A | 0.5
1048 | <Syringe> Syringe failure. | Valve Overload. This error occurs when the valve drive loses steps by blockage or excess back pressure. Continual valve overload errors are an indication the valve should be replaced. | ARMs | Syringes | Error | Rare | N/A | 0.5
1049 | <Syringe> Syringe failure. | Plunger Move Not Allowed. When the valve is in bypass or throughput position plunger movement commands are not allowed. | ARMs | Syringes | Error | Rare | N/A | 0.5
1050 | <Syringe> Syringe failure. | Command Overflow. This error occurs when the command buffer contains too many characters. | ARMs | Syringes | Error | Rare | N/A | 0.5
1100 | Optical reading units voltage control failure. | Voltage out-of-range fault detected. | ORU | Electronic | Error | Rare | N/A | 0.5
1105 | Arm communication error. | RS232 communications error with CAVRO arms | ARMs | Electronic | Error | Rare | N/A | 0.5
1110 | Cuvettes incubator No. 1 temperature control failure. | Incubator temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1111 | Cuvettes incubator No. 2 temperature control failure. | Incubator temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1115 | Optical reading unit No. 1 temperature control failure. | ORU temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1116 | Optical reading unit No. 2 temperature control failure. | ORU temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1117 | Optical reading unit No. 3 temperature control failure. | ORU temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1118 | Optical reading unit No. 4 temperature control failure. | ORU temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1120 | Cuvette shuttle temperature control failure. | Cuvette Shuttle temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1125 | Reagent area temperature control failure. | Reagent Area temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Questionable | Service Call | 1
1130 | Optical reading units data acquisition control failure. | Detector interface fault, e.g., conversion did not complete, emitter light source not correct. | ORU | Electronic | Error | Questionable | Service Call | 1
1131 | Arm <Arm> controller failure | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM | ARMs | Electronic | Error | Rare | N/A | 0.5
1132 | Sample Arm controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1133 | Intermediate Reagent Arm controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1134 | Start Reagent Arm controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1135 | Cuvettes controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | Cuvettes Movement | Cuvette Controller | Error | Rare | N/A | 0.5
1136 | Racks controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | Barcode/Racks | Rack Controller | Error | Rare | N/A | 0.5
1137 | Optical Reading Units controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g., EEPROM) | ORU | Electronic | Error | Rare | N/A | 0.5
1138 | Rack-Cuvette controller failure. | CRC Error detected in writing to or reading from a permanent storage device (e.g. EEPROM) | SW | SW | Error | Rare | N/A | 0.5
1139 | Rack-Cuvette controller failure. | Error detected in updating data on a permanent storage device (e.g. EEPROM) | SW | SW | Error | Rare | N/A | 0.5
1140 | Cuvettes controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | Cuvettes Movement | Cuvette Controller | Error | Rare | N/A | 0.5
1141 | Racks controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | Barcode/Racks | Rack Controller | Error | Rare | N/A | 0.5
1142 | Optical Reading Units controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | ORU | Electronic | Error | Rare | N/A | 0.5
1143 | <Arm> Arm controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | ARMs | Electronic | Error | Rare | N/A | 0.5
1144 | Sample Arm controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1145 | Intermediate Reagent Arm controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1146 | Start Reagent Arm controller failure. | Error detected in updating data on a permanent storage device (e.g., EEPROM) | ARMs | Electronic | N/A | Rare | N/A | 0.5
1151 | <Arm> Arm controller voltage out of range. | Voltage out-of-range fault detected on a controller. | ARMs | Electronic | Error | Rare | N/A | 0.5
1152 | Sample Arm controller voltage out of range. | Voltage out-of-range fault detected on a controller. | ARMs | Electronic | N/A | Rare | N/A | 0.5
1153 | Intermediate Reagent Arm controller voltage out of range. | Voltage out-of-range fault detected on a controller. | ARMs | Electronic | N/A | Rare | N/A | 0.5
1154 | Start Reagent Arm controller voltage out of range. | Voltage out-of-range fault detected on a controller. | ARMs | Electronic | N/A | Rare | N/A | 0.5
1155 | Cuvettes controller voltage out of range. | Voltage out-of-range fault detected on a controller. | Cuvettes Movement | Cuvette Controller | Error | Rare | N/A | 0.5
1156 | Racks controller voltage out of range. | Voltage out-of-range fault detected on a controller. | Barcode/Racks | Rack Controller | Error | Rare | N/A | 0.5
1157 | Optical Reading Units controller voltage out of range. | Voltage out-of-range fault detected on a controller. | ORU | Electronic | Error | Rare | N/A | 0.5
1158 | Rack-Cuvette controller voltage out of range. | Voltage out-of-range fault detected on a controller. | Cuvettes Movement | Electronic | Error | Rare | N/A | 0.5
1160 | Arm <Arm> coordinates error or CTS foot coordinates error. | Automatic coordinates check failure | ARMs | Coordinates | Error | Questionable | Needles/Well Cleaning | 1
1161 | Wash station coordinates error | Wash station coordinates error | ARMs | Coordinates | Error | Questionable | Needles/Well Cleaning | 1
1165 | Arm <Arm> coordinate adjust required. | AM coordinates have never been adjusted using the "Coordinate adjustment" procedure. | ARMs | Coordinates | Error | Questionable | Needles/Well Cleaning | 1
1166 | Missing coordinates of aspiration point in track. | Until the user defines the coordinates of the aspiration point on the LAS track, the instrument is unable to move the LAS probe there. | ARMs | Coordinates | Error | Rare | Needles/Well Cleaning | 0.5
1170 | Invalid system configuration. | Invalid emitter, detection, or incubator module configuration detected, or reference channel not on installed detection module. | SW | SW | Error | Rare | N/A | 0.5
1175 | Sample probe pre-heater temperature control failure. | Probe Pre-Heater temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Rare | N/A | 0.5
1176 | Reagent 1 probe pre-heater temperature control failure. | Probe Pre-Heater temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | N/A | Rare | N/A | 0.5
1177 | Reagent 2 probe pre-heater temperature control failure. | Probe Pre-Heater temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | N/A | Rare | N/A | 0.5
1178 | CTS sample probe pre-heater temperature control failure. | Probe Pre-Heater temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | N/A | Frequent | N/A | 0
1180 | <Probe> probe temperature control failure. | Probe temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Questionable | Service Call | 1
1181 | Reagent 1 probe temperature control failure. | Probe temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Questionable | Service Call | 1
1182 | Reagent 2 probe temperature control failure. | Probe temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | Error | Questionable | Service Call | 1
1183 | CTS sample probe temperature control failure. | Probe temperature defined to be out of range for a period exceeding the fatal error limits. | Temperatures | Temperatures | N/A | Frequent | N/A | 0
1185 | Sample arm clean fluid valve failure. | Valve commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | Error | Rare | N/A | 0.5
1186 | Reagent arm 1 clean fluid valve failure. | Valve commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | N/A | Rare | N/A | 0.5
1187 | Reagent arm 2 clean fluid valve failure. | Valve commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | N/A | Rare | N/A | 0.5
1188 | CTS sample arm clean fluid valve failure. | Valve commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | N/A | Rare | N/A | 0.5
1190 | Sample clean pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | Error | Questionable | Service Call | 1
1191 | Reagent clean pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | Error | Rare | N/A | 0.5
1200 | Sample rinse pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Clean Circuit | Error | Questionable | Service Call | 1
1201 | Reagent arm 1 rinse pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Rinse Circuit | Error | Questionable | Service Call | 1
1202 | Reagent arm 2 rinse pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Rinse Circuit | Error | Questionable | Service Call | 1
1203 | CTS sample arm rinse pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Rinse Circuit | N/A | Rare | N/A | 0.5
1205 | Waste pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Waste Liquid Circuit | Error | Questionable | Service Call | 1
1206 | Coordinate File Saving Failure. | Failure saving coordinates file on the FTP server | SW | SW | Error | Rare | N/A | 0.5
1207 | ORU Stored Reading File Saving Failure. | Failure saving ORU stored reading file on the FTP server | SW | SW | Error | Rare | N/A | 0.5
1208 | CTS sample arm air pump failure. | Pump commanded ON(OFF) but sensor did not indicate ON(OFF) state. | Fluidics | Compressed Air Circuit | Error | Questionable | Service Call | 1
1209 | CTS Failed to Enter Deep wash | The Foot hits the wash station. This would have caused a probe mode switch failure at the end of the wash cycle. | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1210 | Cuvette waste level sensor failure. | Cuvette waste level sensor not responding. | Cuvettes Movement | Cuvette Waste | Error | Questionable | Service Call | 1
1215 | Unable to empty internal waste reservoir. | Waste accumulator full sensor remains on despite cycling of waste pump. | Fluidics | Waste Liquid Circuit | Error | Questionable | Service Call | 1
1220 | Reagent stirring failure. | A check on the power required for stirring returned an error. | Barcode/Racks | Reagent Rack | Error | Questionable | Service Call | 1
1221 | Reagent stirring failure in position <rack position #>, track # <rack track #>. | A check on the power required for stirring position returned an error | Barcode/Racks | Reagent Rack | N/A | Rare | N/A | 0.5
1236 | Cuvette Pivot Arm failure. | Failure on pivot arm movement. | Cuvettes Movement | Loader | Error | Rare | N/A | 0.5
1237 | Cuvette indexer motor failure. | Failure of the indexer motor detected. | Cuvettes Movement | Loader | Error | Rare | N/A | 0.5
1238 | Cuvette Loader Failure. | Failure moving cuvette loader belts detected. | Cuvettes Movement | Loader | Error | Rare | N/A | 0.5
1239 | Cuvette loader failure. | Failure during cuvette indexing detected. | Cuvettes Movement | Loader | Error | Questionable | Remove Cuvette Jam | 1
1240 | Cleaning solution low in clean well. | Validate LLD height in Clean well. If too low, short sipping may occur. | Fluidics | Clean Circuit | Error | Questionable | Service Call | 1
1241 | Probe <arm> Clean priming failure | Clean priming failure | Fluidics | Clean Circuit | Error | Rare | N/A | 0.5
1243 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1244 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1245 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1246 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1247 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1248 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1249 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1250 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1251 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1252 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1253 | Cuvettes shuttle movement failure. | The motor did not complete the requested move. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1254 | Cuvette waste accumulator unable to empty. | Accumulator is full and unable to dump. | Cuvettes Movement | Cuvette Waste | Error | Rare | Remove/Replace Waste Drawer checking Biological Container position. | 0.5
1255 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Questionable | Service Call | 0.25
1256 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Questionable | Service Call | 1
1257 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Questionable | Service Call | 0.25
1258 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Rare | N/A | 0.5
1259 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Questionable | Service Call | 0.25
1260 | Bar code reader movement failure. | The motor did not complete the requested move prior to a motor move time out (bar code). | Barcode/Racks | Barcode | Error | Rare | N/A | 0.5
1261 | All the Optical Reading Units are disabled. Analysis cannot be performed. | All the ORU have been disabled, either manually or automatically. The instrument cannot perform any tests. | ORU | Data Check | Error | Questionable | ORU Calibration | 1
1262 | Flush memory failure. | The requested operation on flush failed | SW | SW | N/A | Rare | N/A | 0.5
1263 | Cuvettes accumulator failure. | The accumulator does not return to the up position within 2 seconds of being emptied. | Cuvettes Movement | Cuvette Waste | Error | Rare | N/A | 0.5
1264 | Optical reading unit No. 1 temperature control failure. | Temperature out of range in a reading unit | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1265 | Optical reading unit No. 2 temperature control failure. | Temperature out of range in a reading unit | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1266 | Optical reading unit No. 3 temperature control failure. | Temperature out of range in a reading unit | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1267 | Optical reading unit No. 4 temperature control failure. | Temperature out of range in a reading unit | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1268 | Cuvettes incubator No. 1 temperature control failure. | Temperature out of range in an incubator | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1269 | Cuvettes incubator No. 2 temperature control failure. | Temperature out of range in an incubator | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1270 | Sample probe temperature or dispense temperature out of range. | Probe temperature or dispense temperature defined to be out of warning range but within failure range. | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1271 | Reagent arm 1 probe temperature or dispense temperature out of range. | Probe temperature dispense temperature defined to be out of warning range but within failure range. | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1272 | Reagent arm 2 probe temperature or dispense temperature out of range. | Probe temperature dispense temperature defined to be out of warning range but within failure range. | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1273 | Sample probe pre-heater temperature out of range. | Probe Pre-Heater temperature defined to be out of warning range but within failure range | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1274 | Reagent arm 1 probe pre-heater temperature out of range. | Probe Pre-Heater temperature defined to be out of warning range but within failure range | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1275 | Reagent arm 2 probe pre-heater temperature out of range. | Probe Pre-Heater temperature defined to be out of warning range but within failure range | Temperatures | Temperatures | Warning | Questionable | Service Call | 1
1276 | Reagent area temperature out of range. | Reagent area temperature defined to be out of warning range but within failure range. | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1277 | Cuvette shuttle temperature out of range. | Cuvette shuttle temperature defined to be out of warning range but within failure range. | Temperatures | Temperatures | Warning | Frequent | N/A | 0
1278 | Optical reading units reference readings out of range for ORU <ORU ID>. | Optical Reference readings not within expected range, possibly due to LED aging or failure. | ORU | Data Check | Warning | Questionable | ORU Calibration | 1
1279 | Optical reading unit <ORU ID> dark readings too high. | ORU dark readings are too high, possibly indicating excessive stray light. | ORU | Data Check | Warning | Questionable | ORU Calibration | 1
1280 | Optical reading units readings too noisy for ORU <ORU ID>. | Optical readings too noisy, possibly due to LED aging or failure. | ORU | Data Check | Warning | Questionable | ORU Calibration | 1
1281 | Optical reading unit <ORU ID> readings out of range. | ORU amplifier and/or A/D converter saturated | ORU | Data Check | Warning | Questionable | ORU Calibration | 1
1282 | Probe <Arm> Rinse Flow Test Failure. | Rinse flow verification procedure failure | Fluidics | Rinse Circuit | Error | Questionable | Check/Prime Rinse circuit; Check Glass Syringes are tighten. | 1
1283 | Optical reading unit raw data calculation error. | Math exception error (e.g., attempt to divide by zero, overflow, result not reasonable). | SW | SW | Warning | Rare | N/A | 0.5
1284 | Air read drift detected for ORU# <ORU ID> Channel# <CHAN ID> Wavelength# <WAVE ID> Drift percentage <DRIFT PERCENTAGE> | Current air read for this wavelength, read head and read channel is outside of the stored air read plus tolerance | ORU | Data Check | Error | Questionable | ORU Calibration | 1
1285 | Probe <probe name> Liquid Level Detection error. | LLD error occurred during pipetting operation. | LLD (Liquid Level Detection) | Missing Liquid | Error | Frequent | N/A | 0
1286 | Probe <probe name> dispense error. | Dispense error. | SW | SW | Error | Frequent | N/A | 0
1292 | Probe <arm> liquid level detection error in LAS track. | LLD error in LAS track. | LLD (Liquid Level Detection) | Missing Liquid | Error | Rare | N/A | 0.5
1293 | Probe <probe name> Liquid Level Detection error in cuvette position <X>, slot <YY>. | LLD error in cuvette | LLD (Liquid Level Detection) | Missing Liquid | Error | Questionable | Needles/Well Cleaning | 0.1
1294 | Probe 1 Liquid Level Detection error in well position 2. | LLD error in well | LLD (Liquid Level Detection) | Missing Liquid | Error | Questionable | Needles/Well Cleaning | 1
1295 | Air read out of range for ORU# <ORU ID> Channel# <CHAN ID> Wavelength# <WAVE ID> Reading: <ORU READING> | Current air read for this wavelength, read head and read channel is out of range | ORU | Data Check | Error | Rare | N/A | 0.5
1296 | Bar code reader invalid message format. | Invalid transmit message format for the bar code reader. | Barcode/Racks | Barcode | N/A | Rare | N/A | 0.5
1297 | RS232 transmit hardware error. | Communications transmission hardware is not functional. | SW | SW | N/A | Rare | N/A | 0.5
1313 | Failed to load cuvette clip at the indexer. | When attempting to move the cuvette indexer to pick up a new cuvette clip, the indexer was not detected to be fully to the left to allow the clip to fall. Move is attempted for five seconds in order to search for the proper position. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1314 | Failed to place cuvette clip in position for pickup. | When the cuvette indexer is attempting to load the cuvette clip into the position for the cuvette shuttle to pick it up, the clip was not detected. Move is attempted for five seconds in order to search for the proper position. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1315 | Pivot arm did not move to the up position. | The move of the pivot arm to the up position was not detected as completed. The motor moves for 10 seconds and the expected sensor to stop the motor was not found. | Cuvettes Movement | Loader | Error | Questionable | Remove Cuvette Jam | 1
1316 | Pivot arm did not move to the down position. | The move of the pivot arm to the down position was not detected as completed. The motor moves for 10 seconds and the expected sensor to stop the motor was not found. | Cuvettes Movement | Loader | Error | Questionable | Remove Cuvette Jam | 1
1317 | Cuvette strip not in shuttle after pick up operation. | After all devices have successfully performed their individual operations to pick up a cuvette, but a cuvette was not detected in the shuttle. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1318 | No cuvette strip in shuttle to drop off. | A request was made to drop off a cuvette, but no cuvette is detected in the shuttle. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1319 | Missing cuvette strip in slot for pick up. | The shuttle is ready to pick up a cuvette, but there is no cuvette in the slot. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1320 | Cuvette shuttle is full. | The shuttle cannot perform a cuvette pick up operation because a cuvette is already detected in the shuttle. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1321 | Cuvette not released into slot. | After all devices have successfully performed their individual operations to drop off a cuvette, but a cuvette was still detected in the shuttle. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1322 | Invalid position to drop off a cuvette. | Improper destination for a cuvette drop off. Cannot drop off a cuvette into the loader for example. | Cuvettes Movement | Shuttle | N/A | Rare | N/A | 0.5
1323 | Cuvette not present to pickup. | The source position for picking up a cuvette is empty. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1324 | Cuvette destination position full. | The destination position for placing a cuvette is occupied. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1325 | Cannot pickup from waste. | The requested cuvette pick up position is the waste position which is not possible. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1326 | Cannot drop off at loader. | The requested drop off position is the loader position which is not possible. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1327 | Cuvette load operation too slow. | Second cuvette move operation attempted prior to the loader/indexer preparing a new cuvette strip. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1328 | Cuvette load operation too slow. | Second cuvette move operation attempted prior to the loader/indexer preparing a new cuvette strip. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1329 | Cuvette slot found too early. | The number of required slots was detected too early in move distance. | Cuvettes Movement | Shuttle | Error | Questionable | Service Call | 1
1330 | Gripper back sensor still active. | During initialization (home), the gripper attempts to move off the back sensor, but the sensor is still detected. OR Extend gripper operation performed, but the back sensor is still detected. | Cuvettes Movement | Shuttle | Error | Questionable | Service Call | 1
1331 | Gripper back sensor not found. | During initialization (home), the gripper moves off the back sensor. Once detected to be off the sensor, the motor moves towards the sensor to make sure it can be detected. If the back sensor is not detected during this search then the alarm is generated. | Cuvettes Movement | Shuttle | Error | Rare | N/A | 0.5
1332 | Gripper forward sensor still active. | After the gripper has completed a retract move the forward sensor is still detected. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1333 | Gripper forward sensor not found. | Extend gripper operation did not detect the front sensor. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
1334 | Cuvette waste tray did not dump. | The accumulator is energized to be in the down position to drop the cuvettes on the tray, but the tray is still detected in the up position. Allowed 250 ms to move. | Cuvettes Movement | Cuvette Waste | Error | Questionable | Remove/Replace Waste Drawer checking Biological Container position. | 1
1335 | Cuvette waste tray not ready to receive cuvettes. | The accumulator is de- energized to go to the up position to get more cuvettes onto the tray, but the tray is still detected in the down position. Allowed 250 ms to move. | Cuvettes Movement | Cuvette Waste | Error | Rare | N/A | 0.5
1336 | Probe <arm> liquid level detection error in LAS arm cover aspiration point. | LLD error in LAS arm cover aspiration point. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1341 | X, Y or Z Axis - Invalid Profile Parameter | X, Y or Z Axis - Invalid Profile Parameter | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1342 | X, Y or Z Axis - Invalid Profile Parameter | X, Y or Z Axis - Invalid Profile Parameter | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1343 | X, Y or Z Axis - Invalid Profile Parameter | X, Y or Z Axis - Invalid Profile Parameter | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1344 | X, Y or Z Axis - Speed Exceeds Profile Maximum | X, Y or Z Axis - Speed Exceeds Profile Maximum | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1345 | X, Y or Z Axis - Speed Exceeds Profile Maximum | X, Y or Z Axis - Speed Exceeds Profile Maximum | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1346 | X, Y or Z Axis - Speed Exceeds Profile Maximum | X, Y or Z Axis - Speed Exceeds Profile Maximum | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1347 | X, Y or Z Axis - Cannot Process Steps Required | X, Y or Z Axis - Cannot Process Steps Required | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1348 | X, Y or Z Axis - Cannot Process Steps Required | X, Y or Z Axis - Cannot Process Steps Required | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1349 | X, Y or Z Axis - Cannot Process Steps Required | X, Y or Z Axis - Cannot Process Steps Required | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1350 | X, Y or Z Axis - Selected Profile Is Invalid | X, Y or Z Axis - Selected Profile Is Invalid | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1351 | X, Y or Z Axis - Selected Profile Is Invalid | X, Y or Z Axis - Selected Profile Is Invalid | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1352 | X, Y or Z Axis - Selected Profile Is Invalid | X, Y or Z Axis - Selected Profile Is Invalid | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1353 | X, Y or Z Axis - Move Took Too Long To Complete | X, Y or Z Axis - Move Took Too Long To Complete | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1354 | X, Y or Z Axis - Move Took Too Long To Complete | X, Y or Z Axis - Move Took Too Long To Complete | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1355 | X, Y or Z Axis - Move Took Too Long To Complete | X, Y or Z Axis - Move Took Too Long To Complete | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1356 | X, Y or Z Axis - Movement Generates Too Many Steps | X, Y or Z Axis - Movement Generates Too Many Steps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1357 | X, Y or Z Axis - Movement Generates Too Many Steps | X, Y or Z Axis - Movement Generates Too Many Steps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1358 | X, Y or Z Axis - Movement Generates Too Many Steps | X, Y or Z Axis - Movement Generates Too Many Steps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1359 | X, Y or Z Axis - Invalid Limit State While Homing | X, Y or Z Axis - Invalid Limit State While Homing | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1360 | X, Y or Z Axis - Invalid Limit State While Homing | X, Y or Z Axis - Invalid Limit State While Homing | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1361 | X, Y or Z Axis - Invalid Limit State While Homing | X, Y or Z Axis - Invalid Limit State While Homing | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1362 | X, Y or Z Axis - Limit Detected During Move | X, Y or Z Axis - Limit Detected During Move | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1363 | X, Y or Z Axis - Limit Detected During Move | X, Y or Z Axis - Limit Detected During Move | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1364 | X, Y or Z Axis - Limit Detected During Move | X, Y or Z Axis - Limit Detected During Move | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1365 | X, Y or Z Axis - Slippage Detected | X, Y or Z Axis - Slippage Detected | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1366 | X, Y or Z Axis - Slippage Detected | X, Y or Z Axis - Slippage Detected | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1367 | X, Y or Z Axis - Slippage Detected | X, Y or Z Axis - Slippage Detected | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1368 | X, Y or Z Axis - Cannot Store More Ramps | X, Y or Z Axis - Cannot Store More Ramps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1369 | X, Y or Z Axis - Cannot Store More Ramps | X, Y or Z Axis - Cannot Store More Ramps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1370 | X, Y or Z Axis - Cannot Store More Ramps | X, Y or Z Axis - Cannot Store More Ramps | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1371 | X, Y or Z Axis - Motor Stalled | X, Y or Z Axis - Motor Stalled | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1372 | X, Y or Z Axis - Motor Stalled | X, Y or Z Axis - Motor Stalled | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1373 | X, Y or Z Axis - Motor Stalled | X, Y or Z Axis - Motor Stalled | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1374 | X, Y or Z Axis - Operation Timeout | X, Y or Z Axis - Operation Timeout | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1375 | X, Y or Z Axis - Operation Timeout | X, Y or Z Axis - Operation Timeout | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1376 | X, Y or Z Axis - Operation Timeout | X, Y or Z Axis - Operation Timeout | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1377 | X, Y or Z Axis - Operation Failed | X, Y or Z Axis - Operation Failed | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1378 | X, Y or Z Axis - Operation Failed | X, Y or Z Axis - Operation Failed | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1379 | X, Y or Z Axis - Operation Failed | X, Y or Z Axis - Operation Failed | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1380 | Tube Release Procedure Failed | Tube Release Procedure Failed | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1381 | Sensor Or Motor Failure Caused Latch Offset Measurement To Fail | Sensor Or Motor Failure Caused Latch Offset Measurement To Fail | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1382 | LLD Hardware Failure | LLD Hardware Failure | LLD (Liquid Level Detection) | Electronic | Error | Rare | N/A | 0.5
1383 | Piercer Was Not in the Expected State | Piercer Was Not in the Expected State | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1384 | Z Axis - Power DAC failed | Z Axis - Power DAC failed | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1385 | Syringe Is Always Busy | Syringe Is Always Busy | ARMs | Syringes | Error | Rare | N/A | 0.5
1386 | Syringe Is Always Busy | Syringe Is Always Busy | ARMs | Syringes | N/A | Rare | N/A | 0.5
1387 | Syringe Took Too Long To Respond | Syringe Took Too Long To Respond | ARMs | Syringes | Error | Rare | N/A | 0.5
1388 | Syringe Could Not Be Initialized | Syringe Could Not Be Initialized | ARMs | Syringes | Error | Rare | N/A | 0.5
1389 | Syringe Does Not Understand The Command | Syringe Does Not Understand The Command | ARMs | Syringes | Error | Rare | N/A | 0.5
1390 | Syringe Command Parameter Is Invalid | Syringe Command Parameter Is Invalid | ARMs | Syringes | Error | Rare | N/A | 0.5
1391 | Syringe Must Be Initialized | Syringe Must Be Initialized | ARMs | Syringes | Error | Rare | N/A | 0.5
1392 | Syringe Plunger Is Overloaded | Syringe Plunger Is Overloaded | ARMs | Syringes | Error | Questionable | Service Call | 1
1393 | Syringe Valve Is Overloaded | Syringe Valve Is Overloaded | ARMs | Syringes | Error | Rare | N/A | 0.5
1394 | Syringe Move Is Not Permitted | Syringe Move Is Not Permitted | ARMs | Syringes | Error | Rare | N/A | 0.5
1395 | Syringe Error With Unknown Source | Syringe Error With Unknown Source | ARMs | Syringes | Error | Rare | N/A | 0.5
1396 | Error Communicating With The Syringe | Error Communicating With The Syringe | ARMs | Syringes | Error | Rare | N/A | 0.5
1397 | Syringe Hardware Malfunction | Syringe Hardware Malfunction | ARMs | Syringes | Error | Rare | N/A | 0.5
1398 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1399 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1400 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1401 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1402 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1403 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1404 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1405 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1406 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1407 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1408 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1409 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1410 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1411 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1412 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1413 | Arm failure. | Arm failure. | ARMs | Universal Arms | N/A | Rare | N/A | 0.5
1414 | CTS piercer lock not locked. | CTS piercer lock not locked. | ARMs | Universal Arms | Error | Rare | N/A | 0.5
1415 | CTS probe mode switch failure. | CTS probe mode switch failure. | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1416 | Cannot switch CTS between piercing and sample mode | The CTS hardware used to change probe between piercing mode and sample mode failed to switch. | ARMs | Universal Arms | Error | Questionable | Service Call | 1
1417 | Location unreachable by <Probe>, track # <rack>, position <Position> | Vial location unreachable | ARMs | SW | Error | Rare | N/A | 0.5
1418 | Location unreachable by <Probe>, track # <rack>, position <Position> | Cuvette location unreachable | ARMs | SW | Error | Questionable | Service Call | 1
1419 | Probe <probe name> unexpected liquid level in rack position <X>, track <YY>. | The probe detects liquid but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1420 | Probe <probe name> unexpected liquid level in cuvette position <X>, slot<YY>. | The probe detects liquid but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1421 | Probe <Probe name> unexpected liquid level in well position <well pos>. | The probe detects liquid but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1422 | Probe <probe name> pre aspiration baseline check failure in rack position <X>, track <YY>. | Invalid liquid detection: difference in baseline values read before and after liquid detection is too low. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1423 | Probe <probe name> pre aspiration baseline check failure in cuvette position <X>, slot <YY>. | Invalid liquid detection: difference in baseline values read before and after liquid detection is too low. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1424 | Probe <probe name> pre aspiration baseline check failure in well position <well pos>. | Invalid liquid detection: difference in baseline values read before and after liquid detection is too low. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1425 | Probe <probe name> aspiration baseline check failure in rack position <X>, track <YY>. | Invalid liquid aspiration: difference in baseline values read before and after aspiration is too high. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1426 | Probe <probe name> aspiration baseline check failure in cuvette position <X>, slot <YY>. | Invalid liquid aspiration: difference in baseline values read before and after aspiration is too high. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1427 | Probe <Probe name> aspiration baseline check failure in well position <well pos>. | Invalid liquid aspiration: difference in baseline values read before and after aspiration is too high. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1428 | Probe <probe name> post aspiration baseline check failure in rack position <X>, track <YY>. | Invalid liquid aspiration. LLD Baseline read after liquid level detection and after liquid aspiration is too high: the probe did not stay submerged. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1429 | Probe <probe name> post aspiration baseline check failure in cuvette position <X>, slot <YY>. | Invalid liquid aspiration. LLD Baseline read after liquid level detection and after liquid aspiration is too high: the probe did not stay submerged. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1430 | Probe <Probe name> post aspiration baseline check failure in well position <well pos>. | Invalid liquid aspiration. LLD Baseline read after liquid level detection and after liquid aspiration is too high: the probe did not stay submerged. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1431 | Probe <arm> unexpected liquid level in LAS track. | The probe detects liquid but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1432 | Probe <arm> pre- aspiration baseline check failure in LAS track. | The probe detects liquid but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1433 | Probe <arm> aspiration baseline check failure in LAS track. | Invalid liquid detection: difference in baseline values read before and after aspiration is too high. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1434 | Probe <arm> post aspiration baseline check failure in LAS track. | Invalid liquid aspiration. LLD baseline read after liquid level detection and after liquid aspiration is too high; the probe did not stay submerged. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1435 | Probe <arm> unexpected liquid level in LAS arm cover aspiration point. | Unexpected liquid level in LAS arm cover aspiration point. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1436 | Probe <arm> pre aspiration baseline check failure in LAS arm cover aspiration point. | Invalid liquid detection: difference in baseline values read before and after liquid detection is too low. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1437 | Probe <arm> aspiration baseline check failure in LAS arm cover aspiration point. | Invalid liquid detection: difference in baseline values read before and after aspiration is too high. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1438 | Probe <arm> post aspiration baseline check failure in LAS arm cover aspiration point. | Invalid liquid aspiration. LLD baseline read after liquid level detection and after liquid aspiration is too high: the probe did not stay submerged. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Critical when >1% of Test Performed | 0.1
1439 | <Probe> probe communications failure. | Cannot communicate with the probe. | ARMs | Electronic | Error | Rare | Service Call | 0.5
1440 | Probe <Probe> unexpected clean level in rack position <X>, track <YY>. | The probe detects clean but the measured position of the liquid is unexpected. | LLD (Liquid Level Detection) | Baseline Error | Error | Questionable | Remove and then replace the Clean | 1
2000 | Rinse solution bottle empty. | The Rinse solution level is below 600 mL. | Fluidics | Rinse Circuit | Error | Frequent | N/A | 0
2005 | Waste tank full. | Waste fluid container Full sensor has been detected. | Fluidics | Waste Liquid Circuit | Error | Frequent | N/A | 0
2006 | Cannot empty waste into container. | Cannot empty waste into container. | Fluidics | Waste Liquid Circuit | Error | Questionable | Empty Tank if Full / Check for Waste Connector Insertion or Tubing obstructions. | 1
2010 | <Cover ID> open. | Safety cover detected open not in diagnostics. | Barcode/Racks | Barcode | Error | Frequent | N/A | 0
2015 | Clean solution bottle empty. | The Clean solution level is below 25 mL. | Fluidics | Clean Circuit | Error | Frequent | N/A | 0
2020 | Cuvette load area empty. | No cuvette strips detected at Empty sensor in the load area. | Cuvettes Movement | Loader | Error | Frequent | N/A | 0
2025 | Cuvettes waste drawer missing. | Cuvette waste drawer sensor detects waste drawer not present. | Cuvettes Movement | Cuvette Waste | Error | Frequent | N/A | 0
2030 | Cuvette waste full | Used cuvettes exceed 80% of the waste drawer capacity. | Cuvettes Movement | Cuvette Waste | Error | Frequent | N/A | 0
2035 | Cuvettes strip drop off failure in position <cuvette slot ID>. | Unable to drop off cuvette in the requested position. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
2040 | Cuvettes strip positioning failure. | The requested cuvette move is not allowed. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
2045 | Cuvette strip pickup failure in position <cuvette slot ID>. | Unable to pick up cuvette in the requested position. | Cuvettes Movement | Shuttle | Error | Questionable | Remove Cuvette Jam | 1
2046 | Clean vial empty. | Clean volume in vial detected to be at error threshold. | Fluidics | Clean Circuit | Error | Rare | Replace bottle if empty / Check for Clean Connector Insertion or Air in circuit. | 0.5
2055 | Insufficient volume in track # <track ID>, position <rack position #>. | Insufficient liquid for the aspiration (sample or reagent). | LLD (Liquid Level Detection) | Missing Liquid | Error | Frequent | N/A | 0
2074 | Insufficient sample volume in LAS track. | Insufficient liquid for the aspirate (LAS track) | LLD (Liquid Level Detection) | Missing Liquid | Error | Frequent | N/A | 0
2075 | Insufficient liquid in cuvette <cuvette slot ID> well <well position>. | Insufficient liquid is detected when mixing is performed in a cuvette cell. | LLD (Liquid Level Detection) | Missing Liquid | Error | Questionable | Service Call | 0.1
2076 | Probe <Arm>: insufficient volume in clean cup. | Insufficient liquid for the aspirate (sample or reagent) | Fluidics | Clean Circuit | Error | Rare | Replace bottle if empty / Check for Clean Connector Insertion or Air in circuit. | 0.5
2077 | Insufficient volume in LAS arm cover aspiration point. | Insufficient liquid for the aspirate (LAS arm cover). | LLD (Liquid Level Detection) | Missing Liquid | Error | Frequent | N/A | 0
2080 | Unrecoverable scheduling time violation. | The actual execution did not meet a critical scheduling time. The job failed and it is not possible to recover | SW | SW | Error | Questionable | HW/SW Error | 1
2081 | Rack in <Track ID> slot unplugged. | A rack already inserted becomes unplugged with the reader not in front of the rack. | Barcode/Racks | Barcode | Error | Frequent | N/A | 0
2082 | ORU Air Calibration data not available for ORU <ORU ID>. | Air Calibration data for the specified ORU are not available. The instrument is not able to perform air reading drift checks. | ORU | Data Check | Error | Questionable | ORU Calibration | 1
2083 | <Material name> insufficient volume. | Volume for the last bottle is detected to be at the error threshold. | LLD (Liquid Level Detection) | Missing Liquid | Error | Frequent | N/A | 0
2084 | An invalid blanking value has been accepted for ORU <ORU ID> (Value = <VALUE>). The ORU will be automatically disabled. | An invalid blanking value has been accepted by the user. The ORU will be automatically disabled. | ORU | Data Check | Error | Rare | N/A | 0.5
2085 | ORU Factor Diluent Calibration data not available for ORU <ORU ID>. | Factor Diluent Calibration data for the specified ORU are not available. The instrument is not able to calculate absolute absorbencies. | ORU | Data Check | Error | Questionable | ORU Calibration | 1
2086 | Maintenance activity interrupted: <Maintenance Activity> | A semi-Automatic or automatic Maintenance Activity's Execution has been interrupted due to an error. | Others | Maintenance | Error | Frequent | N/A | 0
2087 | Maximum number of tests allowed per run exceeded. | There is not more memory in the AM for processing more tests. | SW | SW | Warning | Rare | N/A | 0.5
2091 | CTS Filter missing. | CTS Filter missing. | ARMs | Universal Arms | Error | Questionable | Service Call | 1
2092 | Air presser sensor not operating correctly. | Air presser sensor not operating correctly. | Fluidics | Compressed Air Circuit | Error | Rare | N/A | 0.5
2093 | Not enough free disk space. | System detected not enough free disk space to run the application | SW | SW | Error | Rare | N/A | 0.5
2094 | The default file path for the log files does not exist and cannot be created. | If the default file path does not exist, the system tries to automatically create it. The operation could fail for different causes: - wrong file path specification - insufficient rights to write to the drive - HW/SW failure of the drive | SW | SW | Error | Rare | N/A | 0.5
3001 | <Material name> volume low. | Volume for last bottle detected to be at warning threshold. | LLD (Liquid Level Detection) | Missing Liquid | Warning | Frequent | N/A | 0
3002 | <Material name> on-board stability low. | Remaining stability for the last bottle is detected to be at warning threshold. | Others | Materials | Warning | Frequent | N/A | 0
3003 | <Material name> close to expiration. | Remaining expiration time for the last bottle/lot is detected to be at warning threshold. | Others | Materials | Warning | Frequent | N/A | 0
3004 | Non-identified Material in rack <rack ID> position <rack position #> track # <rack track #>. | A material has been inserted without a bar code ID label. | Others | Materials | Error | Frequent | N/A | 0
3005 | Unknown material in rack <rack ID> position <rack position #> track # <rack track #>. | A material has been inserted with a bar code ID label whose material index does not correspond to any material definition in the material list. | Others | Materials | Error | Frequent | N/A | 0
3006 | New bottle type (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | The bottle type contained in the bar code is different that the bottle type defined in the material definition. | Others | Materials | Warning | Frequent | N/A | 0
3007 | Unknown lot (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been loaded on-board whose lot is not defined in the system. In this case the material ID located in the bar code label is known but the lot ID is not. | Others | Materials | Error | Frequent | N/A | 0
3008 | <Material name> new vial identified in rack <rack ID> position <rack position #> track # <rack track #>. | Difference between the read volume and the current volume is greater than 30%. | Others | Materials | N/A | Frequent | N/A | 0
3009 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3010 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3011 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3012 | Placement error (<material name> in <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3013 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3014 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3015 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3016 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3017 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3018 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3019 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3020 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | A material has been placed in a wrong rack position. | Others | Materials | Error | Frequent | N/A | 0
3021 | Placement error (<material name> in rack <rack ID> position <rack position #> track # <rack track #>). | Two lots of the same material have been placed on-board at the same time. | Others | Materials | Error | Frequent | N/A | 0
3022 | Invalid Demographic Data. | The LIS host has requested a test with invalid demographic data like a birth date greater than current date. | Others | Host | Warning | Frequent | N/A | 0
3023 | Duplicated Rack ID, Rack ID <Rack ID>. | Two different Racks have been identified with the same Rack ID. | Barcode/Racks | Barcode | Warning | Frequent | N/A | 0
3024 | <Test code> <job type> not feasible. Waiting for <material name>. | A job has been sent to the AM and there is not enough volume for some of the reagents or a reagent is not placed. | Others | Materials | Warning | Frequent | N/A | 0
3025 | <Test Code> automatic <job type> not feasible. Waiting for <material name>. | An automatic job created by the AM does not have enough volume for some of the reagents or a reagent is not placed. | Others | Materials | Warning | Frequent | N/A | 0
3026 | QC Due : <TestCode>- <MaterialName> | The expiration time to perform a QC job has entered in its threshold period. | Others | Materials | Warning | Frequent | N/A | 0
3027 | QC Frequency OK: <Test Code>-<MaterialName> | Transition from QC frequency Overdue/Due to OK | Others | QC | Informational | Frequent | N/A | 0
3028 | QC Omitted point: <TestCode>- <MaterialName> | QC result omitted | Others | QC | Informational | Frequent | N/A | 0
3029 | QC Point restored: <TestCode>- <MaterialName> | An omitted QC result has been re-enabled | Others | QC | Informational | Frequent | N/A | 0
3030 | QC Results purged until <LimitDate>: <TestCode>- <Material Name> | QC Results have been purged | Others | QC | Informational | Frequent | N/A | 0
3031 | Message received with incorrect Instrument ID (<Instrument ID>). | The ACL-TOP instrument has received a message from the LIS host with an Instrument ID and/or Host ID, which do not match the configured ones. | Others | Host | Error | Frequent | N/A | 0
3032 | Message received with incorrect Host ID (<Host ID>). | The ACL-TOP instrument has received a message from the LIS host with an Instrument ID and/or Host ID, which do not match the configured ones. | Others | Host | Error | Frequent | N/A | 0
3033 | Invalid Test Code received from host (LIS # <LIS #>). | The LIS host has requested an unknown test. | Others | Host | Warning | Frequent | N/A | 0
3034 | Test orders downloading from host not performed. | The user requests a manual job orders download that the ACL-TOP instrument cannot accomplish. In the same way, the ACL-TOP instrument tries to perform an automatic job orders downloading, and the result is the same. | Others | Host | Error | Frequent | N/A | 0
3035 | Test results uploading not performed. | The user requests a manual test results upload that the ACL-TOP instrument cannot accomplish. In the same way, the ACL-TOP instrument tries to perform an automatic test results uploading, and the result is the same. | Others | Host | Error | Frequent | N/A | 0
3036 | Invalid Test Request received from host (LIS # <LIS #>). | The LIS host has requested a test that is either disabled or inconsistent. | Others | Host | Warning | Frequent | N/A | 0
3037 | Duplicated Test Order received from host (LIS # <LIS #>). | The LIS host has requested a test that is already present in the instrument. | Others | Host | Warning | Frequent | N/A | 0
3038 | Invalid Instrument Sample ID <Instrument Sample ID> received from host. | The LIS host has requested a test for sample whose Instrument Sample ID is unknown. | Others | Host | Error | Frequent | N/A | 0
3039 | LIS request not allowed. | LIS host sent a request that is not implemented for TOP. | Others | Host | Warning | Frequent | N/A | 0
3040 | Invalid Format Message. | The format of the message or fields contained in it are not compliant with the TOP specifications. | Others | Host | Error | Frequent | N/A | 0
3041 | Host Query not performed. | The ACL-TOP instrument tries to perform an automatic host query without succeeding. | Others | Host | Error | Frequent | N/A | 0
3042 | Communications between the instrument and the LIS have been lost. | Communications between the instrument and the LIS have been lost. | Others | Host | Error | Frequent | N/A | 0
3043 | Test Order cancellation not performed (LIS # <LIS #>). | The ACL-TOP could not process a test order cancellation coming from the LIS host. | Others | Host | Warning | Frequent | N/A | 0
3044 | Creation of the LIS communication channel failed. | The ACL-TOP could not create a channel for communicating with the LIS Host. | Others | Host | Error | Frequent | N/A | 0
3045 | Current channel configuration is not supported. | The current channel configuration defined through the control panel is not supported. | Others | Host | Error | Frequent | N/A | 0
3046 | Analyzer Status changed from <AM Status> to <AM Status> | AM status changed | SW | SW | Informational | Frequent | N/A | 0
3047 | Maintenance activity due: <Maintenance Activity>. | An indicated maintenance activity shall be executed by the user. The expiration time to perform the activity has entered in its threshold period. | Others | Maintenance | Warning | Frequent | N/A | 0
3048 | Manual maintenance activity <Maintenance Activity Name> has been added. | A new manual maintenance activity has been defined and added by the user. | Others | Maintenance | Informational | Frequent | N/A | 0
3049 | Manual maintenance activity <Maintenance Activity Name> has been deleted. | A manual maintenance Activity has been deleted by the user. | Others | Maintenance | Informational | Frequent | N/A | 0
3050 | Maintenance activity <Maintenance Activity Name> has been configured. | A maintenance activity has been reconfigured. Their values might be changed. | Others | Maintenance | Informational | Frequent | N/A | 0
3051 | Frequency status for manual maintenance activity <Maintenance Activity Name> has changed to OK. | The frequency status of a maintenance activity has changed to OK. | Others | Maintenance | Informational | Frequent | N/A | 0
3052 | Maintenance activity <Maintenance Activity Name> has been executed. | A maintenance activity has been executed either manually by the user or automatically by the system. | Others | Maintenance | Informational | Frequent | N/A | 0
3053 | Calibration due for <TestCode>. | Calibration frequency in the warning threshold. | Others | Materials | Warning | Frequent | N/A | 0
3054 | Normal Pool Plasma due for <TestCode>. | NPP frequency in the warning threshold. | Others | Materials | Warning | Frequent | N/A | 0
3055 | Analytical Reference due for <TestCode>. | AR frequency in the warning threshold. | Others | Materials | Error | Frequent | N/A | 0
3056 | Calibration of <Test Code> completed: please review and validate. | A calibration has been executed and requires validation from the user. | Others | Materials | Warning | Frequent | N/A | 0
3057 | Calibration of <Test Code> recalculated | A calibration has been recalculated or edited. | Others | Applications | Informational | Frequent | N/A | 0
3058 | Calibration of <Test Code> validated. | A calibration has been manually validated. | Others | Applications | Informational | Frequent | N/A | 0
3059 | Calibration of <Test Code> completed and autovalidated. | A calibration has been completed and autovalidated. | Others | Applications | N/A | Frequent | N/A | 0
3060 | Test Definition <test name> has been added. | A new Test Definition has been added. | Others | Applications | Informational | Frequent | N/A | 0
3061 | Test Definition <test name> has been modified | An existing Test Definition has been modified. | Others | Applications | Informational | Frequent | N/A | 0
3062 | Test Definition <test name> has been deleted. | A Test Definition has been deleted. | Others | Applications | Informational | Frequent | N/A | 0
3063 | Material Definition <material name> has been added. | A new Material Definition has been added. | Others | Materials | Informational | Frequent | N/A | 0
3064 | Material Definition <material name> has been deleted. | A Material Definition has been deleted. | Others | Materials | Informational | Frequent | N/A | 0
3065 | Material Definition <material name> has been modified: ISI value (active Lot). | The Active Lot - ISI value of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3066 | Material Definition <material name> has been modified: on-board stability. | The stability configuration of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3067 | Material Definition <material name> has been modified: assigned value for the Test <test name> (active Lot). | The active Lot - assigned value of a Material Definition & Test Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3068 | Material Definition <material name> has been modified: expiration date (active Lot). | The active Lot - expiration date cfg. of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3069 | Material Definition <material name> has been modified: Lot management | The Lot management cfg. of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3070 | Material Definition <material name> has been modified: Lot Number (active Lot). | The active Lot - Lot Number of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3071 | Material Definition <material name> has been modified: Alternative Lot cfg. | The alternative Lot cfg. of a Material Definition has been changed. | Others | Materials | Informational | Frequent | N/A | 0
3072 | Material Definition <material name> generic change. | An existing Material Definition has been modified. | Others | Materials | Informational | Frequent | N/A | 0
3073 | System Definitions: generic change. | Global change in System Definitions. | Others | Applications | Informational | Frequent | N/A | 0
3074 | System Definitions: Bar code parameters modified. | System Definitions, Bar code parameters changed. | Others | Applications | Informational | Frequent | N/A | 0
3075 | System Definitions: LIS parameters modified | System Definitions, LIS parameters changed. | Others | Host | Informational | Frequent | N/A | 0
3076 | QC Definition <Test name> - <Material name> has been added. | A new QC Definition has been added. | Others | QC | Informational | Frequent | N/A | 0
3077 | QC Definition <Test name> - <Material name> has been deleted. | A QC Definition has been deleted. | Others | QC | Informational | Frequent | N/A | 0
3078 | QC Definition <Test name> - <Material name> has been modified: Target Mean. | The Target Mean of a QC Definition has been changed. | Others | QC | Informational | Frequent | N/A | 0
3079 | QC Definition <Test name> - <Material name> has been modified: Target SD | The Target SD of a QC Definition has been changed. | Others | QC | Informational | Frequent | N/A | 0
3080 | QC Definition <Test name> - <Material name> has been modified: Patient Flags | The Patient Flags of a QC Definition have been changed. | Others | QC | Informational | Frequent | N/A | 0
3081 | QC Definition <Test name> - <Material name> has been modified: Rules Cfg. | The Rules Cfg. of a QC Definition have been changed. | Others | QC | Informational | Frequent | N/A | 0
3082 | QC Definition <Test name> - <Material name> has been modified: Frequency Cfg. | The Frequency cfg. of a QC Definition has been changed. | Others | QC | Informational | Frequent | N/A | 0
3083 | QC Definition <Test name> - <Material name> has been modified: QC Unit | The Unit of a QC Definition has been changed. | Others | QC | Informational | Frequent | N/A | 0
3084 | <Material name> on-board stability expired for last available vial. | Remaining stability for last bottle expired. | Others | Materials | Error | Frequent | N/A | 0
3085 | <Material name> expired. | The present date is greater than the expiration date of the last available bottle/lot. | Others | Materials | Error | Frequent | N/A | 0
3086 | <Material name> requires stirring (rack <rack ID> position <rack position #> track # <rack track #>). | A material requiring stirring is placed in a position not having stirring capability. | Others | Materials | Error | Frequent | N/A | 0
3087 | QC Overdue: <TestCode>- <MaterialName> | The expiration date for a QC has been reached. | Others | Materials | Error | Frequent | N/A | 0
3088 | QC Failed: <TestCode>- <MaterialName> | QC rules have failed. It is a hint to detect problems in the measurement. | Others | Materials | Error | Frequent | N/A | 0
3089 | Maintenance activity overdue: <Maintenance Activity>. | An indicated Maintenance Activity should already have been executed by the user. The expiration date for a Maintenance Activity has been reached. | Others | Maintenance | Warning | Frequent | N/A | 0
3090 | Calibration Overdue for <TestCode>. | Frequency expired for the active validated calibration. | Others | Materials | Error | Frequent | N/A | 0
3091 | Normal Pool Plasma Overdue for <TestCode>. | The NPP frequency has expired. | Others | Materials | Error | Frequent | N/A | 0
3092 | Analytical Reference Overdue for <TestCode>. | The AR frequency has expired. | Others | Materials | N/A | Frequent | N/A | 0
3093 | Analytical Reference Failed for <TestCode>. | Analytical Reference has failed | Others | Materials | N/A | Frequent | N/A | 0
3094 | Calibration for <Test Code> interrupted. | A Calibration was active and it was manually or automatically interrupted. | Others | Materials | Warning | Frequent | N/A | 0
3095 | Auto validation of completed calibration for <Test Code> failed. | A calibration has been completed and the auto validation failed. | Others | Materials | Warning | Frequent | N/A | 0
3096 | <Test Code> is inconsistent and was removed from <Rack ID>, position <rack position #>. | Programmed test was not consistent and was removed. | Others | Materials | Warning | Frequent | N/A | 0
3097 | <TestCode> <JobType> job cannot be executed and was removed from sample/material <Sample ID/Material Name>. Number of predilution cuvette slots is insufficient. | Number of predilution cuvette slots is insufficient to execute a calibration or a parallelism job. | Others | Materials | Warning | Frequent | N/A | 0
3098 | <TestCode> <JobType> job not feasible. Calibration not available. | A patient job has been sent to AM and there is no validated calibration for the on board lots. | Others | Materials | Warning | Frequent | N/A | 0
3099 | <TestCode> <JobType> job not feasible. QC failed for <Material Name>. | A patient job has been sent to AM and there is a QC failed for that test. | Others | Materials | Warning | Frequent | N/A | 0
3100 | Software access permissions have been modified. | The user changed some access levels in the Software access screen. | Others | Applications | Informational | Frequent | N/A | 0
3101 | User Definition <User Name> has been added. | A new User Definition has been added. | Others | Applications | Informational | Frequent | N/A | 0
3102 | User Definition <User Name> has been deleted. | A new User Definition has been added. | Others | Applications | Informational | Frequent | N/A | 0
3103 | User Definition <User Name> has been modified. | An existing User Definition has been modified. | Others | Applications | Informational | Frequent | N/A | 0
3104 | User Definition <User Name> has been modified> Security Level | The security Level of an existing User Definition has been changed. | Others | Applications | Informational | Frequent | N/A | 0
3105 | User Definition <User Name> has been locked by the system. | The user intended to log in with erroneous password 3 consecutive times, so the system has locked this user. | Others | Applications | Informational | Frequent | N/A | 0
3106 | Successful log in by user: <user>. | User was recognized and could log in to the system. | Others | Applications | Informational | Frequent | N/A | 0
3107 | Unsuccessful log in by user: <user>. | User could not be recognized and could not log in to the system. | Others | Applications | Informational | Frequent | N/A | 0
3108 | Log out by user: <user>. | User logged out from the system. | Others | Applications | Informational | Frequent | N/A | 0
3109 | The user modified patient demographics for sample <sample ID>. | The user modified patient demographic information. | Others | Applications | Informational | Frequent | N/A | 0
3110 | Reflex Rule <Reflex Rule Name> has been added. | A new Reflex Rule has been added. | Others | Applications | Informational | Frequent | N/A | 0
3111 | Reflex Rule <Reflex Rule Name> has been deleted. | An existing Reflex Rule has been deleted. | Others | Applications | Informational | Frequent | N/A | 0
3112 | Reflex Rule <Reflex Rule Name> has been modified. | Reflex Rule has been modified. | Others | Applications | Informational | Frequent | N/A | 0
3120 | Auto-Validation Rules have been modified. | Auto-Validation Rules have been modified. | Others | Applications | Informational | Frequent | N/A | 0
3130 | Sender Definition <Sender Code> has been added. | A new Sender Definition has been added. | Others | Applications | Informational | Frequent | N/A | 0
3131 | Sender Definition <Sender Code> has been deleted. | An existing Sender Definition has been deleted. | Others | Applications | N/A | Frequent | N/A | 0
3132 | Sender Definition <Sender Code> has been modified. | An existing Sender Definition has been modified. | Others | Applications | Informational | Frequent | N/A | 0
3140 | Test Profile <Test Profile Name> has been added. | A new Test Profile Definition has been added. | Others | Applications | Informational | Frequent | N/A | 0
3141 | Test Profile <Test Profile Name> has been deleted. | A Test Profile Definition has been deleted. | Others | Applications | Informational | Frequent | N/A | 0
3142 | Test Profile <Test Profile Name> has been modified. | A Test Profile Definition has been modified. | Others | Applications | Informational | Frequent | N/A | 0
3145 | QC Profile <QC Profile Name> has been added. | A new QC Profile Definition has been added. | Others | QC | Informational | Frequent | N/A | 0
3146 | QC Profile <QC Profile Name> has been deleted. | A QC Profile Definition has been deleted. | Others | QC | Informational | Frequent | N/A | 0
3147 | QC Profile <QC Profile Name> has been modified. | A QC Profile Definition has been modified. | Others | QC | N/A | Frequent | N/A | 0
3161 | Test Definition <source test Name> has been promoted to <target test Name> | A Test Definition has been promoted. | Others | Applications | N/A | Frequent | N/A | 0
3164 | Vial Properties <Vial Type Name> <Vial Properties Revision> has been added. | A Vial Properties has been added by the user. | Others | Applications | N/A | Frequent | N/A | 0
3165 | Vial Properties <Vial Type Name> <Vial Properties Revision> has been updated. | A Vial Properties has been modified by the user. | Others | Applications | Informational | Frequent | N/A | 0
3166 | Vial Properties <Vial Type Name> <Vial Properties Revision> has been deleted. | A Vial Properties has been deleted by the user. | Others | Applications | N/A | Frequent | N/A | 0
3167 | Vial Set <Vial Set Name> has been added. | A Vial Set has been added by the user. | Others | Applications | N/A | Frequent | N/A | 0
3168 | Vial Set <Vial Set Name> has been updated. | A Vial Set has been updated by the user. | Others | Applications | Informational | Frequent | N/A | 0
3169 | Vial Set <Vial Set Name> has been deleted. | A Vial Set has been deleted by the user. | Others | Applications | N/A | Frequent | N/A | 0
3180 | LAS parameters modified. | LAS definitions parameters changed. | Others | Applications | N/A | Frequent | N/A | 0
3181 | IM parameters modified. | IM parameters changed. | Others | Applications | N/A | Frequent | N/A | 0
3182 | <Test code> is inconsistent/disabled. | There was a programmed test for an LAS sample. The test definition is inconsistent/disabled and the test was removed. | Others | Applications | Warning | Frequent | N/A | 0
3183 | Sample volume for <test code> exceeds the limit of a cuvette cell. | The LAS sample volume required to complete a single test cannot exceed 535 μL. | Others | Applications | Warning | Frequent | N/A | 0
3184 | <Test code> <job type> job not feasible. Sample <sample ID> has expired. | A job cannot start because the LAS sample expired. | Others | Applications | Error | Frequent | N/A | 0
3185 | <Test code> <Job type> job not feasible. Waiting for sample <sample ID>. | A job has been sent to the AM and there is not enough sample volume. | Others | Applications | Warning | Frequent | N/A | 0
3186 | Imported Params Report files deleted. Analytical Version in files is different than current one. | Files to generate Imported Params report have been deleted. | Others | Applications | Warning | Frequent | N/A | 0
3187 | <TestCode> <JobType> Required Clean Material shortage has been estimated. Monitor <material name> availability. | A job has been sent to AM and there is not enough volume for a required clean material or the clean is not placed. | Others | Materials | Warning | Frequent | N/A | 0
3200 | Storage of messages sent to LIS <%> full. | The UDC storage is almost full of messages waiting to be sent to LIS. | Others | Host | Warning | Frequent | N/A | 0
3201 | Storage of messages sent to LIS overloaded. New uploaded messages will be rejected. | The UDC storage is almost full of messages waiting to be sent to LIS. No more messages are accepted. | Others | Host | Error | Frequent | N/A | 0
3202 | Upload message was rejected and therefore not sent to LIS. | UDC Rejected Uploaded Message. | Others | Host | Error | Frequent | N/A | 0
3211 | Auto Start failed | Auto Start could not be started. | SW | SW | Warning | Frequent | N/A | 0
3212 | <Controller Name> controller has been upgraded to version <Version Name>. | Controller upgraded successfully. | SW | SW | N/A | Rare | N/A | 0.5
3213 | The system detected a software version designed for experimental use. | Software version is for experimental use only. A controller, master, IM or CM is compiled in a debug or the version name is experimental. CM is a demo installation. | SW | SW | Warning | Rare | N/A | 0.5
3214 | Interface module has been upgraded to version <version name>. | IM has been upgraded. | SW | SW | N/A | Rare | N/A | 0.5
3215 | LAS sample expiration time exceeded for <data dictionary - sample ID>. | The time that an LAS sample can remain placed has been exceeded. | Others | Materials | Error | Frequent | N/A | 0
3216 | Semiautomatic maintenance activity <Maintenance Activity Name> has been executed. | A maintenance activity has been executed either manually by the user or automatically by the system. | Others | Maintenance | Informational | Frequent | N/A | 0
3217 | Automatic maintenance activity <Maintenance Activity Name> has been executed. | A maintenance activity has been executed either manually by the user or automatically by the system. | Others | Maintenance | Informational | Frequent | N/A | 0
4000 | Unable to move bar code reader. | The bar code reader cannot move because the rack is not fully inserted or removed. | Barcode/Racks | Barcode | Warning | Questionable | Rack removed too fast or rack labels damaged. Place in and out again slowly the rack. | 0.1
4004 | Invalid bar code size | Bar code size bigger than the maximum size of sample ID (15 characters) | Barcode/Racks | Barcode | Error | Rare | N/A | 0.5
4005 | Rack rejected. | The bar code label read is invalid. | Barcode/Racks | Barcode | Warning | Questionable | Rack removed too fast or rack labels damaged. Place in and out again slowly the rack. | 0.3
4006 | CTS rack rejected. | CTS rack rejected. | Barcode/Racks | Barcode | Warning | Frequent | N/A | 0
4007 | Cap not detected for sample in rack <Rack ID>, position <position #>, track <Track ID>. | Cap not detected by the piercer. | ARMs | Universal Arms | Error | Frequent | N/A | 0
4015 | Rack button identification failure. | The rack button pressed is not a valid code. | Barcode/Racks | Barcode | Warning | Rare | N/A | 0.5
4020 | Bar code reader communication failure. | Cannot communicate with the bar code reader. | Barcode/Racks | Barcode | Warning | Questionable | Service Call | 1
4025 | In-line cleaning solution level low. | The in-line cleaning solution level is below 75 mL. | Fluidics | Clean Circuit | Warning | Frequent | N/A | 0
4030 | Rinse solution level low. | Rinse solution level is below 1000 mL. | Fluidics | Rinse Circuit | Warning | Frequent | N/A | 0
4035 | Cuvettes low in the load area. | No cuvette strips detected at Warning sensor in load area. | Cuvettes Movement | Loader | Warning | Frequent | N/A | 0
4040 | Cuvette waste almost full. | The level of cuvettes in the waste drawer is close to maximum height. | Cuvettes Movement | Cuvette Waste | Warning | Frequent | N/A | 0
4050 | Waste tank almost full. | Waste fluid container Warning sensor detected. | Fluidics | Waste Liquid Circuit | Warning | Frequent | N/A | 0
4051 | A new set of ORU air blanking values has been saved. | A new set of ORU air blanking values has been saved. | ORU | Data Check | Informational | Frequent | N/A | 0
4052 | A new set of ORU factor diluent blanking values has been saved. | A new set of ORU factor diluent blanking values has been saved. | ORU | Data Check | Informational | Frequent | N/A | 0
4053 | Probe <Arm>: Coordinates adjustment procedure completed. | Probe Coordinates adjustment procedure completed. | ARMs | Coordinates | Informational | Frequent | N/A | 0
4054 | The analyzer is unable to process the request to enable CTS piercing probe mode. | The analyzer is unable to process the request to enable CTS piercing probe mode. | ARMs | Universal Arms | Error | Rare | N/A | 0.5
4055 | The analyzer is unable to process the request to disable CTS piercing probe mode. | The analyzer is unable to process the request to disable CTS piercing probe mode. | ARMs | Universal Arms | Error | Rare | N/A | 0.5
4056 | Waste tank not present. | Waste tank is not present. | Fluidics | Waste Liquid Circuit | Warning | Frequent | N/A | 0
4057 | CM in use is from a different ACL TOP model. | The system detected the instrument model is different than during the last connection. | SW | SW | Warning | Rare | N/A | 0.5
4058 | Unable to create raw data report <path&filename>. | Unable to create raw data report file. | SW | SW | Warning | Rare | N/A | 0.5
4059 | LAS Interface Manager (IM) communications error. | LAS Interface Manager (IM) communications error. | Others | Track Communication | Error | Frequent | N/A | 0
4060 | Test counter statistics have been reset. | User reset the test counter statistics | Others | Applications | Informational | Frequent | N/A | 0
4061 | A new software version has been created with following label: <Data Dictionary-SW Version>. | A new system software version has been detected. | Others | Applications | Informational | Frequent | N/A | 0
4062 | A new hardware event has been created with following label: <Data Dictionary-HW Event>. | A new Hardware event has been added by the user. | Others | Applications | Informational | Frequent | N/A | 0
4063 | Description: <Description> Functional Area: <Functional Area> | User entered a manual alarm. | Others | Applications | Informational | Frequent | N/A | 0
5000 | LLD error | The AM has detected liquid where liquid should not be (a false positive) | LLD (Liquid Level Detection) | Baseline Error | Error | Rare | Critical when >1% of Test Performed | 0.5
5002 | Error detected in ORU | Generic ORU error. If the signal check for a particular test is disabled, all ORU flags are treated as an error. | ORU | Electronic | Error | Rare | N/A | 0.5
5090 | (Prep) Material dispense error | AM reported a material dispense error. An unexpected material dispense error was detected by the AM software. Not expected behavior. | SW | SW | Error | Rare | N/A | 0.5
5100 | Measured result failed | The result unit is the measured result and DR has been unable to compute it | Others | Applications | Warning | Frequent | N/A | 0
5190 | Temperature out of range | Generic flag indicating that at least one temperature reading was out of range. All temperature out of range alarms are mapped to a single flag within each determination. | Temperatures | Temperatures | Warning | Rare | N/A | 0.5
5191 | Warning detected in ORU. | Generic ORU warning. If signal checks are enabled in the test definitions for the particular test, all ORU flags are treated as warning. | ORU | Electronic | Warning | Rare | N/A | 0.5
5192 | Material stirring warning | The AM has detected that the material stirring mechanism is not stirring the material and this is a coagulation test (stirring prolongs material stability) | Barcode/Racks | Reagent Rack | Warning | Rare | N/A | 0.5
5568 | Calibration and parallelism transformations do not match | Calibration transformations do not match parallelism transformations. | Others | Applications | Warning | Frequent | N/A | 0
5650 | Material stability | One or more material's stability has expired for this determination. | Others | Materials | Warning | Frequent | N/A | 0
5651 | Material expired | One or more of the materials used to run this determination has expired. | Others | Materials | Warning | Frequent | N/A | 0
5660 | LAS sample expired during processing. | LAS sample expired during processing. | Others | Materials | Warning | Frequent | N/A | 0
5700 | Volume tracking - insufficient liquid (sample) | Error indicating that there was insufficient sample to pipette. | Others | Materials | Error | Frequent | N/A | 0
5701 | Volume tracking - insufficient liquid (reagent) | The AM determined that there was insufficient reagent to pipette. | Others | Materials | Error | Frequent | N/A | 0
5702 | Missing post dispense rinse/clean detected. | Error indicating missing dispense of rinse or clean was detected in the software due to timing issues in the scheduler. | Fluidics | Rinse Circuit | Error | Rare | N/A | 0.5
5703 | Material dispense error | Because of a scheduling error, a material that should have been dispensed into the reaction cuvette was not dispensed. This check is done for safety purposes and it is intended that this error will never be seen. | SW | SW | Error | Rare | N/A | 0.5
05751 | SampleTube Underfilled | The instrument detected sample tube is underfilled during first sample aspiration. Low Limit TFH + Low Limit calibrated - 0.7mm |  |  | Warning
05760 | Fluidic Obstruction | The instrument detected a sample aspiration occlusion. |  |  | Warning
05770 | Hemoglobin High | Sample concentration of hemoglobin too high and could affect test results. |  |  | Warning
05771 | Bilirubin (Icterus) High | Sample concentration of bilirubin too high and could affect test results. |  |  | Warning
05772 | Limpemia High | Sample Lipemia too high and could affect test results. |  |  | Warning
05773 | HIL Test Pending | Temporary flag to indicatethe HIL test is stillpending executions and calculation. When HIL result arrives, it will be replaced. |  |  | Warning
05774 | HIL Test Failed | Unable to calculate HIL at sample or job level. |  |  | Warning
05775 | HIL Test Cancelled | HIL test cancelled by the user. The system will not request more HIL data. |  |  | Warning
05776 | HIL not available | HIL results not available: HIL is not feasible because there is not enough sample. |  |  | Warning
05777 | HIL out of range | End point algorithm for at least one wavelength for HIL is greater than threshold. |  |  | Warning
05851 | Sample tube undefined | The instrument detected sample tube is underfilled during first sample aspiration. Low Limit TFH + Low Limit calibrated - 0.7mm |  |  | Error
05860 | Fluidic Obstruction | The instrument detected a sample aspiration occlusion. |  |  | Error
05870 | Hemoglobin High | Sample concentration of hemoglobin too high and could affect test results. |  |  | Error
05871 | Bilirubin (Icterus) High | Sample concentration of bilirubin too high and could affect test results. |  |  | Error
05872 | Lipemia High | Sample Lipemia too high and could affect test results. |  |  | Error
10000 | LAS track communications error. | LAS track communications error. | Others | Track Communication | Error | Frequent | N/A | 0
10002 | LAS track unavailable. | LAS track unavailable. | Others | Track Communication | N/A | Frequent | N/A | 0
10003 | LAS sample <sample ID> rejected. | LAS track has provided sample data with inconsistencies. | Others | Track Communication | Error | Frequent | N/A | 0
10004 | LAS sample <sample ID> rejected. | Capped tube provided by the LAS track. | Others | Track Communication | Error | Frequent | N/A | 0
10005 | Reset queue request denied. Allow to clear the physical queue is unsafe. | LAS track has requested to reset the queue but the instrument does not let it do that. | Others | Track Communication | Error | Frequent | N/A | 0
10006 | LAS sample <sample ID> cannot be processed at this time. | The instrument rejects the LAS sample because it is not ready for sampling. | Others | Track Communication | Warning | Frequent | N/A | 0
10007 | LAS sample <sample ID> rejected. | LAS sample was intended to be provided while a reset queue was in progress. | Others | Track Communication | Warning | Frequent | N/A | 0
10008 | LAS sample <sample ID> not expected at aspiration point. This sample shall not be processed. | Sample not expected at aspiration point. | Others | Track Communication | Error | Frequent | N/A | 0
10009 | LAS sample <sample ID> order different than expected at aspiration point. This sample shall be processed. | Sample not expected at aspiration point. | Others | Track Communication | Warning | Frequent | N/A | 0
10010 | LAS sample <sample ID> not processed. | Processing not feasible for LAS sample <sample ID>. Sample has either -No jobs ordered -Jobs that are inconsistent/disabled -Required volume exceeds the maximum cuvette volume. | Others | Track Communication | Warning | Frequent | N/A | 0
10011 | LAS sample <sample ID> not processed. | Sample not processed due to sample tube queue empty. | Others | Track Communication | Warning | Frequent | N/A | 0
10012 | LAS sample <sample ID> not processed. | Sample not processed due to a fluid not detected. | Others | Track Communication | Error | Frequent | N/A | 0
10013 | LAS sample <sample ID> not processed. | Sample not processed due to an insufficient volume. | Others | Track Communication | Warning | Frequent | N/A | 0
10014 | LAS sample <sample ID> not processed. | Sample not processed due to an LAS arm failure. | Others | Track Communication | Error | Frequent | N/A | 0
10015 | LAS sample <sample ID> not processed. | Sample ID mismatch when releasing the sample tube. | Others | Track Communication | Error | Frequent | N/A | 0
10016 | LAS track request not supported. | LAS track sent a request that is not implemented/ supported for TOP. | Others | Track Communication | Error | Frequent | N/A | 0
10017 | Invalid format message. | The format of the message or fields contained in it are not compliant with the TOP specifications. | Others | Track Communication | Error | Frequent | N/A | 0
10018 | Creation of the LAS track communication channel failed. | The ACL-TOP could not create a channel for communicating with the LAS track. | Others | Track Communication | Error | Frequent | N/A | 0
10019 | Current channel configuration is not supported. | The current channel configuration defined through the control panel is not supported. | Others | Track Communication | Error | Frequent | N/A | 0
10020 | Storage of messages sent to the LAS track <%> full. | The Synapse storage is almost full of messages waiting to be sent to the LAS track. | Others | Track Communication | Warning | Frequent | N/A | 0
10021 | Storage of messages sent to the LAS track overloaded. New uploaded messages will be rejected. | The Synapse storage is almost full of messages waiting to be sent to the LAS track. No more messages are accepted. | Others | Track Communication | Error | Frequent | N/A | 0
10022 | LAS track communications error. | Message not delivered to the LAS track. | Others | Track Communication | Error | Frequent | N/A | 0
10023 | LAS sample <sample ID> rejected. | Tube type provided by the LAS track not supported. | Others | Track Communication | Error | Frequent | N/A | 0
10024 | LAS sample <sample ID> not processed. | Sample not processed due to an instrument hardware error. | Others | Track Communication | Error | Frequent | N/A | 0
10025 | Test orders received for LAS sample <sample ID>, which is not located in the queue. | Sample not found in queue after test order from LIS. | Others | Track Communication | Warning | Frequent | N/A | 0
10026 | IM software error. | IM Software internal exception. | Others | Track Communication | Error | Frequent | N/A | 0
10027 | LAS track communications error. | Message not responded to by the LAS track. | Others | Track Communication | Error | Frequent | N/A | 0
10028 | LAS sample <sample ID> not supported. This tube shall be released. | Sample ID not supported. | Others | Track Communication | Error | Frequent | N/A | 0
10029 | LAS sample <sample ID> not processed. | Sample not processed due to LAS sample aspiration timeout. | Others | Track Communication | N/A | Frequent | N/A | 0